Failure Mode and Effects Analysis (FMEA) Risk Analysis Example

Acupuncture Needles

| Process Step | Failure Mode | Effect | Sev 1-3 | Cause | Controls | Prob 1-3 | Risk Index | Recommended Actions |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Inspection | Defective needle is missed | Patient injury - bruise | 2 | Insufficient sampling plan |  | 1 | 1 |  |
|  |  |  |  | Operator error (missed defect) |  | 2 | 4 | Provide training and visual guidance on inspection techniques |
|  |  | Patient injury – broken needle under skin | 3 | Insufficient sampling plan |  | 1 | 3 |  |
|  |  |  |  | Operator error (missed defect) |  | 2 | 6 | Provide training and visual guidance on inspection techniques |
|  |  | Unusable needle | 1 | Insufficient sampling plan |  | 1 | 1 |  |
|  |  |  |  | Operator error (missed defect) |  | 2 | 2 | Provide training and visual guidance on inspection techniques |
| Load autoclave | Load is not balanced | Incomplete sterilization | 3 | Operator Error | Document states to distribute load evenly | 2 | 6 | Add load diagram to document |
|  |  |  |  |  | Bioburden QC testing | 1 | 3 | Validate bioburden test method |
|  | Unit is over-loaded | Incomplete sterilization | 3 | Operator Error | Document provides load/program matrix | 2 | 6 | Validate load capacities |
|  |  |  |  |  | Bioburden QC testing | 1 | 3 | Validate bioburden test method |
|  | Unit is under-loaded | Deformed needle housing – unusable needle | 1 | Operator Error | Document provides load/program matrix | 2 | 2 | Validate load capacity |
|  |  |  |  |  | Visual inspection during packaging/ assembly operation | 2 | 2 |  |
| Run autoclave program | Incorrect program is selected | Incomplete sterilization | 3 | Operator Error | Lot file review includes equipment printout review | 2 | 6 | Fix batch sizes and use a single program |
|  |  |  |  |  | Bioburden QC testing | 1 | 3 | Validate bioburden test method |
|  | Entire cycle is not run | Incomplete sterilization | 3 | Equipment failure | Preventive maintenance | 1 | 3 |  |
|  |  |  | 3 | Power failure |  | 2 | 6 | Connect to emergency generator; allow repeat run |
|  | Cycle does not meet max temperature | Incomplete sterilization | 3 | Equipment is out of calibration | Annual calibration check | 2 | 6 | Run each cycle with biological indicators |

# Text boxes

Inspect needles

Load needles into autoclave bags

Unload autoclave

Transfer bags to packaging

Run program

Load autoclave